Full-Stokes Spectral Reconstruction for a Channeled Spectropolarimeter

- Project Overview

This repository provides four algorithms for full-Stokes spectral reconstruction in channeled spectropolarimetry (CSP):

CCSP: A constrained reconstruction algorithm based on a sparse dictionary of Legendre polynomials and the discrete cosine transform (DCT).

DDNN: A spectral reconstruction algorithm based on a data-driven neural network.

PGNN: A physics-driven neural network reconstruction algorithm incorporating a polarization modulation physical model.

DA-PGNN: A two-stage reconstruction algorithm combining data-driven pretraining and physics-driven fine-tuning.

- Algorithm Descriptions

  - CCSP

CCSP constructs a sparse dictionary using Legendre polynomials and discrete cosine transform (DCT) basis functions, and reconstructs Stokes spectra by solving a constrained optimization problem.

This method uses MATLAB and the CVX toolbox for optimization and does not rely on neural-network training. The reconstruction process takes the measured modulation spectrum as input and recovers the spectrum by combining the polarization modulation physical model with sparse representation.

  - DDNN

DDNN is a data-driven neural-network method that establishes a mapping between measured modulation spectra and Stokes spectra through supervised learning.

During training, data containing modulation spectra and their corresponding Stokes-spectrum labels are used. During inference, the measured modulation spectrum is input directly, and the reconstructed Stokes spectrum is output.

The first stage is trained using a loss function centered on the forward reconstruction error of the modulation spectrum.

  - PGNN

PGNN incorporates a polarization modulation physical model into the neural-network training process to improve the consistency between the reconstruction results and the actual optical system.

The network directly receives the measured modulation spectrum and predicts the corresponding Stokes spectrum. The prediction is then passed through the physical forward model to calculate the predicted modulation spectrum. A physics-constrained loss is constructed by comparing the predicted modulation spectrum with the measured modulation spectrum.

Depending on the implementation, spectral smoothness constraints and degree-of-polarization constraints can also be introduced to improve the stability and physical plausibility of the reconstruction results.

  - DA-PGNN

DA-PGNN adopts a two-stage strategy that combines data-driven pretraining with physics-driven fine-tuning.

Stage 1: Data-driven pretraining. Training data with Stokes-spectrum labels are used to learn an initial mapping from modulation spectra to Stokes spectra.

Stage 2: Physics-driven fine-tuning. Starting from the pretrained network, the polarization modulation physical model and corresponding regularization constraints are introduced to further optimize the reconstruction results.

By combining prior information provided by the training data with constraints provided by the physical model, this method helps improve the reconstruction quality of complex polarization spectra.

- Runtime Environment

  - Python Environment

DDNN, PGNN, and DA-PGNN are implemented using Python and TensorFlow.

  - MATLAB Environment

The CCSP algorithm requires the following software:

- MATLAB

- CVX convex optimization toolbox

Before running the algorithm, install and configure CVX according to the official CVX instructions, and ensure that CVX can correctly call the required solver.

- Input Data Format

  - Measured Modulation Spectrum

During neural-network inference, the measured modulation spectrum is read directly. A typical CSV file contains two columns:

| Column Name | Meaning |
| --- | --- |
| Wavelength (nm) | Wavelength in nanometers (nm) |
| I | Measured modulation-spectrum intensity |

  - Supervised Training Data

The data-driven pretraining stages of DDNN and DA-PGNN require labeled training data. Whether PGNN requires Stokes labels depends on its specific training implementation; training with purely physical constraints does not require ground-truth Stokes labels.

The training-data file contains six columns:

| Column | Data |
| --- | --- |
| 1 | Wavelength (nm) |
| 2 | S0 |
| 3 | S1 |
| 4 | S2 |
| 5 | S3 |
| 6 | I |

- Physical Model Parameter Configuration

PGNN and DA-PGNN use a polarization modulation physical model to construct forward constraints.

Before running the algorithms, check and configure the relevant parameters according to the experimentally calibrated system, including:

- The refractive-index difference or phase-delay parameters of the retarders;

- The thicknesses of the retarders;

- The fast-axis orientations of the two retarders;

- The transmission-axis orientation of the analyzer;

- Intensity normalization and scaling parameters for the modulation spectrum and Stokes parameters.

Do not treat example parameters as calibration results for the actual experimental system without verification.The physical model maps the Stokes spectra predicted by the network to the corresponding modulation spectra and calculates the physics-constrained loss. This does not mean that ground-truth Stokes spectra are used to generate the input data at the reconstruction stage.

If the optical configuration of the actual instrument differs from the PSIM model assumed in the code, the forward model should be modified accordingly.

- Notes

Before running the algorithms, check the data paths, wavelength sampling, and intensity normalization settings.

The physical-model parameters used by PGNN and DA-PGNN should match the actual experimental system.

The supervised pretraining stages of DDNN and DA-PGNN require training data with Stokes labels.